Specifications for  Electrochemical Workstation with Impedance Spectroscopy system and RDE/RRDE attachment

No. of channels required: 	+2 Independent channels.(The system should work as 2 independent channels as well as Bi-Potentiostat mode )

Current ranges for channel 1 & 2: 	±10 nA to ±500 mA or more

Compliance voltage: 			±12 V or better

Applied Potential range:		 ±10 V or more

Voltage Resolution: 			1 μV or better

Current Resolution: 			1 pA or better

Potentiostat Rise time: 			<500ns or better

Frequency range: 	10 μHz-6 MHz or more – 1 channel should have EIS facility

AC Amplitude:				0.5 mV to 2 V

Acquisition speed:			 700,000 samples/second or better

Bandwidth of electrometer:		 6 MHz or better

Electrode configurations: 		 2, 3, 4 & 5 electrode configurations.

Interface:					 USB

Computer:	I-5 processor, >4 GB RAM, with sufficient USB ports,  Windows -8, 500 GB HDD, 22” screen

Accessories required:

RDE/ RRDE Specifications:

Rotar/Rotating speed:	Electrode rotar with CE and ETL marks with RPM between 50-10,000 RPM with <10 RPM resolution

RDE Electrode:				5.0 mm GC- 5 No’s

5.0 mm Platinum -1 No.

5.0 mm gold -1 No.

5.0 mm silver -1 No.

RRDE electrode:	Pt. Ring & GC Disc-1 No. Necessary cell system for RDE/RRDE.

Pt. Ring  & Gold Disc-1 No.

Cell for RDE/RRDE setup*:	150 ml with water jacket, Stand, lid with sleeve, stoppers

Hg/Hgo Reference Electrode*:	1 No.

Ag/Agcl Reference Electrode*:	1 No.

Platinum counter electrode*:		1 No.

Glassy carbon Electrode:			3 No.

Floating facility is required.

Software’s required.

Battery/Super capacitor testing software including Impedance measurement facility, constant load discharge, constant current constant power etc..

Voltammetry software

Fuel cell/photovoltaic testing software

Pulse software including DPV,NPV, RNPV,SWV etc

Corrosion software  including LPR, Tafel etc.

Equivalent fitting circuit software

Warranty: 01  Year

(*All accessories must be compatible with RDE/RRDE setup)

Quotations should be addressed to the below address on or before 18.10.16

Prof. P. Selvam

Dept. of Chemistry

IITM , 600 036

Ph:04422574235/4200