|  | Department of Aerospace Engineering Indian Institute of Technology Madras Chennai – 600 036, India |
| --- | --- |
| Dr. M.S. Manikandan Project Coordinator |  |

Ref: No. ASE/MSM/ LASER/ ISRO /2015/      					DATE: 05.06.2015

Dear Sir, 								                 Due date :  26.06.2015

- Quotations are invited in duplicate for the various items shown below/overleaf/enclosed list.

- The quotations are to be in two parts as Technical Offer and as Commercial offer:

The two parts of the offer are to be clearly marked on the envelopes. The two parts of the offer in separate envelopes must be enclosed in the one bigger envelope duly sealed and super scribed with reference number and due date and must be addressed to the undersigned so as to reach him on or before the due date stipulated above.

- Fax and Email quotation are not acceptable.

- Quotations should be valid for 60 days from the due date and period of delivery required , warranty terms etc. should also be clearly indicated. A minimum of one year warranty is required from the date of commissioning.

- Imported supplies should be quoted for CIF Madras.

- Local firms to quote for free delivery to this Institute. If quoted for Ex-Godown delivery charges be indicated separately.

- Relevant literature pertaining to the items quoted with full specifications (and drawing, if any) should be sent along with the Quotations, wherever applicable. Samples / machine/ equipment if called for should be submitted / demonstrated at free of charges, and collected back at the supplier’s expenses. Compliancy certificate is to be provided indicating conformity to the technical specifications

- Sales Tax/General Taxes/ED if applicable and such other taxes legally leviable and intended to be claimed should be distinctly shown along with the price quoted. If this is not indicated no such claim will be admitted at any stage. The taxes leviable should take into consideration that we are entitled to have concessional Sales Tax applicable to Non-Government Educational Institutions run with no profit motive for which a concession is given. Sales Tax Certificate will be issued at the time of final settlement of the bill.

- Goods should be supplied carriage paid and insured.

- Goods shall not be supplied without an official supply order.

- If the item is under DGS&D Rate contract No. and the price must be mentioned.  It may also please be indicated whether the supply can be made direct to us at the Rate contract price (Please note that we are not Direct Demanding Officers).  If so please send copy of the RC.

- The Guarantee period of the item may be indicated clearly.

- In case of LC. Payment, 90% of the payment will be made after completion of the supply. The balance 10% of the payment will be made after satisfactory installation of the equipment.

- IIT Madras is exempt from payment of Excise Duty and is eligible for concessional rate of custom duty. Necessary certificate will be issued on demand. IIT Madras will make necessary arrangements for the clearance of imported goods at the Airport/Seaport. Hence the price should not include the above charges.

- Acceptance and Rejection:- I.I.T. Madras has the right to accept the whole or any part of the Tender or portion of the quantity offered or reject it in  full without assigning any reason.

Yours faithfully,

Project Coordinator

Items required: Continuous wave (CW), modulated DPSS/OPS Laser as per specifications enclosed.

--------------------------------------------------------------------------------------------------------------------------------------------

Phone Nos.:  (044) 2257 4025 / 5000 FAX: (044) 2257 4002, E-mail: manims@ae.iitm.ac.in, manims@iitm.ac.in

Continuous wave (CW), modulated DPSS/OPS Laser

Technical specifications required to be met by the vendors are given below:

Application: Fluid flow-diagnostics

Output power:5W or larger (should be able to set the power at any value between 0 and the maximum value)

Laser CLASS:4

Wavelength:532 nm (green laser)

Beam quality (M2):Lesser than 2.0

Beam diameter: 4mm or lesser

Beam divergence:Lesser than 2.0mrad

Polarization ratio: 100:1 or better

Output power stability:Lesser than 10% over 4 hours

Modulation: The laser must have the capability to operate at asingle pulse mode that can be modulated up to10 kHz. Double pulse mode of operation should also be possible.

Spatial mode: Near-TEM00 or TEM00

Operation lifetime hours:More than 10000 hours

Operating temperature: 10-350C

Operating voltage: 85-260 VAC

Warm-up time:Lesser than 15 minutes

Additional Requirements:

- Vendors should provide continuous technical support and maintenance of equipment.

- Vendors have to provide warranty for a minimum of one year.

- Vendors must have sufficient experience in supplying equipment in reputed organizations for research purpose. They must provide references of end users whom we can contact for their experience with the supplied machine. Experience of the end users will also be used as a criterion for the selection of bids that meet technical requirements.

- Vendors must provide detailed documentation for the equipment.

- Vendors may be called to visit and give presentation/demonstration on the equipment after opening the technical bid. They need to provide the approximate date for this presentation in the bid. The time period for this presentation would be 14 days from the date of opening of the bid.

- Vendors must provide training to our technical staff for using the equipment.

- All the expenses for installation, training and post sales technical support will be borne by the vendor.