|  | Department of Aerospace Engineering Indian Institute of Technology Madras Chennai – 600 036, India |
| --- | --- |

K. BHASKAR

Head of the Department

Ref: No. ASE/SRCH/ KEROSENE MASS FLOW CONTROLLERS /2013/                            DATE : 1 .10. 2013

DUE DATE :   21 .10 .2013

Dear Sir,

- Quotations are invited in duplicate for the various items shown below/overleaf/enclosed list.

- The quotations are to be in two parts as

Technical Offer and as Commercial offer:

The two parts of the offer are to be clearly marked on the envelopes. The two parts of the offer in separate envelopes must be enclosed in the one bigger envelope duly sealed and super scribed with reference number and due date and must be addressed to the undersigned so as to reach him on or before the due date stipulated above.

- The quotations duly sealed and super scribed on the envelope with reference no. and due date, should be addressed to the undersigned so as to reach him or before the due date stipulated above. Fax and Email quotation are not acceptable.

- Quotations should be valid for 60 days from the due date and period of delivery required , warranty terms etc. should also be clearly indicated. A minimum of one year warranty is required from the date of commissioning.

- Imported supplies should be quoted for CIF Madras.

- Local firms to quote for free delivery to this Institute. If quoted for Ex-Godown delivery charges be indicated separately.

- Relevant literature pertaining to the items quoted with full specifications (and drawing, if any) should be sent along with the Quotations, wherever applicable. Samples / machine/ equipment if called for should be submitted / demonstrated at free of charges, and collected back at the supplier’s expenses. Compliancy certificate is to be provided indicating conformity to the technical specifications

- Sales Tax/General Taxes/ED if applicable and such other taxes legally leviable and intended to be claimed should be distinctly shown along with the price quoted. If this is not indicated no such claim will be admitted at any stage.The taxes leviable should take into consideration that we are entitled to have concessional Sales Tax applicable to Non-Government Educational Institutions run with no profit motive for which a concession is given. Sales Tax Certificate will be issued at the time of final settlement of the bill.

- Goods should be supplied carriage paid and insured.

- Goods shall not be supplied without an official supply order.

- If the item is under DGS&D Rate contract No. and the price must be mentioned.  It may also please be indicated whether the supply can be made direct to us at the Rate contract price (Please note that we are not Direct Demanding Officers).  If so please send copy of the RC.

- In case of LC. Payment,90% of the payment will be made after completion of the supply. The balance 10% of the payment will be made after satisfactory installation of the equipment.

- IIT Madras is exempt from payment of Excise Duty and is eligible for concessional rate of custom duty. Necessary certificate will be issued on demand. IIT Madras will make necessary arrangements for the clearance of imported goods at the Airport/Seaport. Hence the price should not include the above charges.

- Acceptance and Rejection:- I.I.T. Madras has the right to accept the whole or any part of the Tender or portion of the quantity offered or reject it in  full without assigning any reason.

Yours faithfully

HEAD OF THE DEPARTMENT

Items required: Kerosene Mass Flow Controllers as per specifications enclosed. Quantity required: 2 Nos.

--------------------------------------------------------------------------------------------------------------------------------------------

Phone Nos.:  (044) 2257 5000 / 4012  FAX: (044) 2257 4022,  E-mail : sujith@ iitm.ac.in

| 2/2 | Department of Aerospace Engineering Indian Institute of Technology Madras Chennai – 600 036, India |
| --- | --- |

Technical specification of Kerosene mass flow controller

The mass flow controller is used to maintain a specified fuel flow rate to the combustors used for studying transverse combustion instabilities.

Brief scope of supply

- Mass flow controller containing flow meter, control valve, control unit etc. in order to maintain a specified fuel flow rate shall be supplied.

- The specification of the two mass flow controllers is tabulated below.

| S. no | 1 | 2 |
| --- | --- | --- |
| Medium | kerosene | kerosene |
| Flow rate | 0.5-5.5LPM | 0.1-1LPM |
| Accuracy | <±0.5% | <±0.5% |
| Output | 4-20mA | 4-20mA |
| Wetted parts | Kerosene compatibility | Kerosene compatibility |
| Operating temperature | Room temperature | Room temperature |
| Quantity | 1 no | 1 no |

- Flow indicators of quantity 2, shall be supplied with an input for 4-20mA and the required transmitter power supply. The indicator shall be compatible for computer interface.

- Documents like operating manual, trouble shooting, regular maintenance of equipment etc. shall be provided.

- The functionality of the mass flow controller and accessories in total to the stipulated specification shall be demonstrated to an IITM representative after installation.

- The warranty certificate for the equipment shall be provided for a period of 12 months.