Tender Specifications for Gas Chromatograph (GC)

| Chromatographic Performance |
| --- |
| Retention time performance | Retention time repeatability < 0.06% or < 0.006min Area repeatability < 2% RSD |
| System Capabilities |
| GC Configuration | Must support simultaneously: - One Capillary S/SL inlet with EPC or PPC - FID detector with EPC or PPC |
| Column Oven |
| Operating temperature range | Oven operating range must be from Ambient temperature +8 °C to 425 °C or better |
| Temperature set point resolution | Oven temperature set point resolution must be 1 °C or better |
| Oven ramps/plateaus | Oven must support 5 oven ramps with 6 plateaus or better |
| Oven ramp rates | Maximum achievable Temperature ramp rates: 75 °C/min (120 V units are limited to 30oC/min) |
| Oven cool down | Oven cool down (22 °C ambient) 450 to 50 °C in 8.0 min |
| Maximum run time | Maximum run time: 999.99 min (16.7 h). |
| Ambienttemperature rejection | Oven must have ambient rejection: < 0.01 °C per 1 °C |
| Oven power safety | Oven power must turn-off automatically when the lid/door is opened. |
| Electronic PneumaticControl |
| Electronic PneumaticControl | Pneumatics must be electronic and programmable through inlets to detectors. |
| Pressure setpoint precision and control | Pressure set points must be able to be adjusted by increments of 0.01 psi or better |
| Flow sensors | Flow sensors must have specification for Accuracy: <± 5% depending on carrier gas |
| Constant flow mode | Constant flow mode must be available |
| Split Splitless Inlet (SSL) |
| Split Splitless Inlet (SSL) | Inlet must have Electronic pneumatic control of carrier, split and septum purge gases, including electronic ON/OFF |
| Capillary column id SSL | Must be suitable for all capillary columns (50 µm to 530 µm id) |
| Maximum temperature SSL | Maximum temperature: 400 °C. |
| EPC Pressure Ranges SSL | EPC available in 0 to 60 psig (0 to 680 kPa) |
| Total flow range SSL | Must be able to set total flow range: 0 to 200 mL/min N2 0 to 500 mL/min H2 or He |
| Inlet maintenance SSL | Inlet sealing system is built in standard with each S/SL inlet for quick, easy, injector liner changes in under 30 seconds. |
| Flame IonizationDetector (FID) |
| Sensitivity FID | Minimum detectable level (for tridecane): < 3 pg C/s |
| Linear Dynamic Range FID | Linear dynamic range: >107 (± 10%). |
| Linear Dynamic Range FID | Full-range digital data path enables peaks to be quantified over the entire 107 concentration range in a single run, without the need to set a switch or software setting to obtain this full range of data. |
| Data rate FID | Must be able to set data rate up to 100 Hz |
| EPC control FID | Must have electronic pneumatic control of make-up, H2 and air flows (with electronic ON/OFF). -Air: 0 to 800 mL/min -H2 : 0 to 100 mL/min -Makeup gas (N2 or He): 0 to 100 mL/min in two selectable modes: constant flow or column flow plus makeup flow constant |
| Necessary accessories, Gas cylinders, Filters & PC & Printer should be offered. |  |