|  | Department of Aerospace Engineering Indian Institute of Technology Madras Chennai – 600 036, India |
| --- | --- |

Prof. K. BHASKAR

HEAD OF THE DEPARTMENT

Ref: No. ASE/AK/ BOMB CALORIMETER / 2014/    			DATE: 10.02.2015

Due Date : 02.03.2015

Dear Sir,

- Quotations are invited in duplicate for the various items shown below/overleaf/enclosed list.

- The quotations are to be in two parts as Technical Offer and as Commercial offer:

The two parts of the offer are to be clearly marked on the envelopes. The two parts of the offer in separate envelopes must be enclosed in the one bigger envelope duly sealed and super scribed with reference number and due date and must be addressed to the undersigned so as to reach him on or before the due date stipulated above.

- Fax and Email quotation are not acceptable.

- Quotations should be valid for 60 days from the due date and period of delivery required, warranty terms etc. should also be clearly indicated. A minimum of one year warranty is required from the date of commissioning.

- Imported supplies should be quoted for CIF Madras.

- Local firms to quote for free delivery to this Institute. If quoted for Ex-Godown delivery charges be indicated separately.

- Relevant literature pertaining to the items quoted with full specifications (and drawing, if any) should be sent along with the Quotations, wherever applicable. Samples / machine/ equipment if called for should be submitted / demonstrated at free of charges, and collected back at the supplier’s expenses. Compliancy certificate is to be provided indicating conformity to the technical specifications

- Sales Tax/General Taxes/ED if applicable and such other taxes legally leviable and intended to be claimed should be distinctly shown along with the price quoted. If this is not indicated no such claim will be admitted at any stage. The taxes leviable should take into consideration that we are entitled to have concessional Sales Tax applicable to Non-Government Educational Institutions run with no profit motive for which a concession is given. Sales Tax Certificate will be issued at the time of final settlement of the bill.

- Goods should be supplied carriage paid and insured.

- Goods shall not be supplied without an official supply order.

- If the item is under DGS&D Rate contract No. and the price must be mentioned.  It may also please be indicated whether the supply can be made direct to us at the Rate contract price (Please note that we are not Direct Demanding Officers).  If so please send copy of the RC.

- The Guarantee period of the item may be indicated clearly.

- In case of LC. Payment, 90% of the payment will be made after completion of the supply. The balance 10% of the payment will be made after satisfactory installation of the equipment.

- IIT Madras is exempt from payment of Excise Duty and is eligible for concessional rate of custom duty. Necessary certificate will be issued on demand. IIT Madras will make necessary arrangements for the clearance of imported goods at the Airport/Seaport. Hence the price should not include the above charges.

- Acceptance and Rejection:- I.I.T. Madras has the right to accept the whole or any part of the Tender or portion of the quantity offered or reject it in  full without assigning any reason.

Yours faithfully,

- sd-

HEAD OF THE DEPARTMENT

Items required: BOMB CALORIMETER as per specifications enclosed.

-----------------------------------------------------------------------------------------------------------------------------------

Phone Nos.: (044) 2257 4001/ 4019/5000    FAX: (044) 2257 4002    E-mail:amitk@ae.iitm.ac.in

Specifications of Bomb Calorimeter

Measurement variables: heat of combustion, gross calorific value of solid, liquid fuels and propellants

Scope of application: hydrocarbon liquid fuels,solid fuels, cellulosic fuels, biomass, plastics, propellants (example: HTPB + ammonium perchlorate, nitro-glycerine + nitrocellulose), laboratory education

Analysis type: isoperibol, adiabatic or dynamic that allows measurements from equipment to conform to ASTM D240-02, D4809-00, E144-94, D5865 and other equivalent methods

Measurement units: J/kg, cal/gm, BTU/lb

Temperature resolution: 0.0001°C or better

Combustion Bomb: Halogen and acidresistant stabilized stainless steel

Resolution: 0.001 kcal/gm or better

Relative standard deviation: 0.1% or better

Measurement range: up to 40,000 J/gm

Crucible type: corrosion resistant alloy or quartz

The equipment should also be compact table top, operable standalone (without PC and software).The equipment should have a display (touch screen is preferred) of results in standalone operation (Optional: upgradeable to a software controlled system).In stand-alone mode the equipment should be operable manually (for education) or in automatic mode. In automatic mode calculation and display of calorific values should be automatic.In-built data storage capability for previous analyses and provision for transferring data to PC or media is preferred.

Oxygen/gas filling and water filling may be manual, semi-automatic or automatic. The unit may be with or without chiller unit

Analysis time per sample: 15 to 30 minutes or lower

Power supply: 230 V, 50/60 Hz

Three to five references of users of the equipment to be provided for feedback.

Consumables and Accessories

O2 gas pressure regulator with filter cartridge

O2 gas cylinder

Benzoic acid pellets (at least 50nos)

Sample crucibles (quartz preferred or corrosion resistant alloy steel, 2 nos)

Gelatine capsules (100 nos.)

Ignition threads for at least 1000 analysis

Ignition fuse wire(5 nos)

Bomb O-rings (5 nos)

Calibration standard with certificates.

Warranty 1 year at the minimum from the date of successful installation

Optional (Consumables and Accessories):

Additional halogen and acid resistant stabilized stainless steel combustion bomb

Additional normal combustion bomb

Additional crucibles (quartz or corrosion resistant alloy steel, 2nos)

Calorimeter software if sold separately

Chiller if sold separately

Additionalconsumables:ignition thread (1000 tests), benzoic acid pellets (50 nos), O-rings (5 nos), ignition fuse wire (5 nos), gelatine capsules (5 nos).

Annual maintenance cost (AMC) for additional two years beyond warranty period